Ethanol Dilution Method for D-Series Bio-Oil: Carbon Overstatement Assessment

To: Isometric

From: Max Lavine, Carbon Protocol & Verification Lead, Charm Industrial

Date: October 13, 2026

Question

Does the ethanol dilution method overstate the carbon content of D-series bio-oil, and therefore CO₂e stored?

Charm’s Position

No. Across four paired samples, ethanol-corrected carbon averaged 4.8 percentage points below the undiluted result. At the 84th-percentile bound applied in Isometric Standard v2.0 §3.7.2, the difference remains negative (−1.9 points).

Relation to PDD

“The Bio-Oil is sampled for CHN testing by an ISO accredited laboratory after the completion of Sparging. The % C reported by the laboratory is multiplied by the Mass of Bio-Oil Injected. That value is multiplied by 3.67 or 44/12 to convert the carbon content to CO2e.” (PDD v1.12, Section B, Gross CO2e Sequestration)

The dilution method changes two steps: laboratory sample preparation (absolute ethanol added at a weighed fraction) and derivation of the reported % C (blend result corrected for that fraction). The sampling point, laboratory CHN analysis, Mass of Bio-Oil Injected and the 44/12 conversion are unchanged.

1. Paired Results (ASTM D5291, SGS St. Rose)

| Sample | Undiluted C, % (m/m) | Diluted C, % (m/m), ethanol-corrected | Difference (points) | SGS report |
| --- | --- | --- | --- | --- |
| 9-2026-W1-DC | 42.5 | 35.7 | −6.8 | [report no. to confirm] |
| 9-2026-W3-DC | 41.6 | 42.3 | +0.7 | [report no. to confirm] |
| 9-2026-W4-DC | 35.3 | 24.6 | −10.7 | SR26-05820.002; SR26-05820.001 Reissue 1 |
| 10-2026-W1-DC | 35.7 | 33.2 | −2.5 | SR26-05962.001; .002 |
| Mean | 38.8 | 34.0 | −4.8 |  |

Diluted samples were blended at 15.0 wt% (9-2026-W4) and 14.9 wt% (10-2026-W1) absolute ethanol. Carbon was corrected run by run using 52.14 % C for ethanol. SGS’s dilution worksheet reports, for each run, the weighed ethanol fraction, measured blend carbon and corrected carbon.

2. Upper Bound on Overstatement

Overstatement would appear as a positive difference (diluted − undiluted). The question is how large that difference could plausibly be, given the four pairs. Isometric’s threshold for a conservative value:

“Isometric recommends that the value used in the Removal or Reduction calculation be either ≤16th or ≥84th percentile, depending on which one yields a more Conservative estimate of Removal or Reduction.” (Isometric Standard v2.0 §3.7.2)

| Statistic (difference = diluted − undiluted, n = 4) | Value (points of C) |
| --- | --- |
| Mean difference (standard error) | −4.8 (2.5) |
| 84th-percentile upper bound | −1.9 |

Diluted results were lower in 3 of 4 pairs. The 84th-percentile bound is the mean difference plus the 84th-percentile t-value for three degrees of freedom times the standard error (−4.8 + 1.19 × 2.5 = −1.9). The bound is below zero: at the percentile §3.7.2 sets for a Conservative estimate, ethanol-corrected carbon does not exceed undiluted carbon.

3. Accuracy Against a Homogeneous Reference Oil

Charm’s validation on AECN oil, which can be measured neat, blended ethanol at 33.38 wt%:

“The average carbon content for each method, Figure 2, differs by only 0.22%. Analysis via t-test showed that the 95% confidence interval was -0.338% to +0.778%.” (Justification for MRV analysis of D-series oil via ethanol dilution, “Validation of CHN Analysis With Ethanol Dilution”)

Neat AECN averaged 42.62 % C and diluted AECN 42.40 % C, so the interval is for neat minus diluted. The largest overstatement by the diluted method consistent with that interval is 0.34 points.

4. Repeatability on D-Series Oil

| Sample | Undiluted runs, range (n = 10) | Ethanol-corrected runs, range (n = 3) |
| --- | --- | --- |
| 9-2026-W4-DC | 19.8 – 50.6 % C | 23.5 – 25.6 % C |
| 10-2026-W1-DC | 20.6 – 48.8 % C | 30.9 – 34.7 % C |

Undiluted runs on the same sample divide between a low-carbon and a high-carbon group, consistent with each 0.1 g subsample drawing mainly one phase. SGS noted on both 9-2026-W4-DC reports: “Please Note - Samples Are Predominantly Water.” Ethanol evaporation between blending and analysis would lower the corrected value, by 0.20–0.32 points per wt% of ethanol lost.

Requested Confirmation

Charm requests that Isometric confirm the ethanol dilution method is acceptable for CHN determination of D-series bio-oil for MRV.